## Slide 1: Progress Report on the Study Design theme
William Vogt, Ph.D.
U.S. Food and Drug Administration
DISCLAIMER: The mention of commercial products, their sources, or their use in connection with material reported herein is not to be construed as either an actual or implied endorsement of such products by the Department of Health and Human Services. The opinions described herein are solely those of the author and should not be considered FDA policy or recommendations.

## Slide 2: Overview of Standards
What is a “Standard”?
  Test object (phantom)
  Test method (image quality testing)
  Data format (DICOM)
Documented set of technical definitions, guidelines
  Specific to a type of device, system, process
  Based on stakeholder consensus
What does a standard do?
  Define performance characteristics
  Describe performance test methods
  Specify performance thresholds, safety requirements

## Slide 3: Levels of Standardization
Test Methodology Research

## Slide 4: Levels of Standardization
Collaboration, Consensus Documents
Test Methodology Research

## Slide 5: Levels of Standardization
Collaboration, Consensus Documents
Standards Development Orgs.
Test Methodology Research

## Slide 6: Levels of Standardization
Collaboration, Consensus Documents
Standards Development Orgs.
Accreditation (ACR,MQSA)
Test Methodology Research

## Slide 7: Why Develop Standards for PAI?
Nascent technology with wide variety of designs, applications
Standardized test methods are needed to support:
  Objective, quantitative performance evaluation
  Design evaluation and optimization
  Interdevice comparison (benchmarking)
  Regulatory decision-making
  QA/QC, marketing, promotion
  Indication of technological maturity
FDA Medical Device Development Tools (MDDT) Program
  FDA can review and qualify tools, methods as suitable for regulatory purposes
  Qualified for specific “context of use”, e.g., photoacoustic image quality assessment
  Can be led/submitted by internal or external tool developers
  Reviewed by team of FDA experts

## Slide 8: Elements of Medical Device Standards
Scope
  As broad as possible…
  Without sacrificing rigor, utility
Example: Which types of PAI devices?
  Handheld vs. annular vs. hemispherical
  Clinical devices vs. preclinical research tools
A generalized set of recommendations
  May be established based on current type(s)…
  But readily adapted to other or future types

## Slide 9: Study Design Theme
Overall goal: Design multi-centre studies and resolve logistical challenges.
Aims:
Develop pre-specified study protocol for PAI system characterization
Resolve multi-site logistics (shipping, phantom calibration/constancy testing, etc.)
Why do a multi-site phantom study?
  Reduce phantom-based test methods to practice
  Identify topics where consensus must be reached
  Produce peer-reviewed publications referenced during standard development

## Slide 10: Previous Activity
Initial teleconferences held in mid/late 2018
Summary of initial recommendations
Identification of new consensus challenges

## Slide 11: Previous Activity
Multi-site phantom study: how will phantom stability be evaluated?
Fabrication by a single source site
  Measure optical + acoustic properties, mass
  Maintain reference phantom, plugs
  Cambridge group
“Test-ship-test” protocol
  Source site characterizes phantom before shipment
  Test site performs imaging
  Source site reevaluates phantom
  Ship to next test site
Source Site
Test Site
Test Site
Test Site
Test Site

## Slide 12: Previous Activity
Initial consensus challenge: to what degree should standard prescribe measurement techniques?
  Devices may have conflicting instructions for optimal use
  Lack of controlled guidelines lowers reproducibility, consistency of the test method
Establish general guidelines
  “Lock in” acquisition and display settings for entire test barrage
  Protocol prescribes acoustic couplants (water, gel)
  Positioning control
    Stages: specify scan spacing (e.g., < half acoustic wavelength for elevational resolution testing)
    Handheld devices: alignment procedure to produce “best-case scenario”
  User training on phantom handling, packaging

## Slide 13: Previous Activity
Image quality analysis should be pre-specified
  Which performance characteristics, definitions?
  Statistical measurements (multiple phantom locations)?
  ROI analysis (ROI size, number of targets within FOV)
Image quality metrics
  Axial, lateral, elevational resolution
  Penetration depth, sensitivity
  Intensity uniformity
  Spatial measurement accuracy
  Low-contrast detectability
  Spectral measurement accuracy (SO2, contrast agent conc.)
  Artifacts
Feedback into phantom design (Phantom Theme)

## Slide 14: Ongoing Activity
How should study results be reported to source site?
Standardize a written report with specified structure/sections
  Device design parameters (wavelength, frequency, beam geometry)
  Alignment procedure
  General descriptions of processing (reconstruction, spectral unmixing, averaging, filtering)
  Display settings (dynamic range, TGC, color map)
  Quantitative test results
How will data be shared between sites? (DAM Theme)

## Slide 15: Ongoing Activity
We want to develop a robust, consensus-based study protocol
Several identified challenges to address first
A consensus list of terms and definitions is needed
  Improve clarity of communication in IPASC activities
  Identify and resolve potential points of ambiguity
  Harmonize compliance with existing IEC/ISO standards
May 2019: Initial draft circulated for comment
July: Collated comments, prepared revised document
August: Follow-up t-con with commenters
November: Final draft circulated for adoption

## Slide 16: Consensus List of Terms and Definitions
List of normative references (other standards)
  ANSI Z136.1:2014: American National Standard for Safe Use of Lasers
  IEC 61391-1:2006: Ultrasonics – Pulse-echo scanners – Part 1: Techniques for calibrating spatial measurement systems and measurement of system point-spread function response
  IEC 61391-2:2006: Ultrasonics – Pulse-echo scanners – Part 2: Measurement of maximum depth of penetration and local dynamic range
  ISO 80000-7:2008: Quantities and units – part 7: Light
Avoid re-inventing the wheel

## Slide 17: Consensus List of Terms and Definitions
Radiometry terms
  Spectral bandwidth, pulse duration, pulse repetition rate
  Harmonization with existing standardized terminology
  Exception: PRR, not PRF (“frequency” used in acoustic terms)
An early issue: what does everybody mean by “fluence”?
  Fluence = energy incident from all directions on a small point
  Fluence is used to describe subsurface distribution of energy
  Many papers use fluence = incident energy / beam area on tissue
  Incident energy per area is to be called “radiant exposure”
Sliney, Photochemistry and Photobiology 83, 2007

## Slide 18: Consensus List of Terms and Definitions
Why should we care about clarifying this distinction?
  Consistent with ANSI, ISO, and IUPAC usage
  ANSI laser safety limits are defined based on radiant exposure
    Fluence is tissue-dependent, hard to measure
    Fluence can exceed radiant exposure by ~5x
  In PAI, we work with both incident and subsurface quantities
  “incident fluence” could be R.E. or fluence in shallow tissue voxel
IPASC plans to adopt use of “radiant exposure” for the quantity “incident pulse energy per beam area on tissue”
  “Radiant exposure” (correct)
  “Incident fluence” (ambiguous)
  “Fluence” (incorrect)
Sliney, Photochemistry and Photobiology 83, 2007

## Slide 19: Consensus List of Terms and Definitions
Transducer parameters
  Spatial coordinates (axial, lateral, elevational axes)
  Element/array properties (e.g., pitch, element sensitivity)
  Frequency response (center frequency, fractional bandwidth)
Tissue properties
  What does it mean for a phantom material to be “tissue-mimicking material”?
    Definition defines general characteristics, level of description of an acceptable TMM
    Material whose optical/acoustic properties are similar to those of a declared tissue type over specified optical wavelength and acoustic frequency ranges
    List of characterized properties “including but not limited to…”
Image quality metrics
  Not necessarily a final recommended set of metrics to use

## Slide 20: Consensus List of Terms and Definitions
One key new concept: minimum detectable optical absorption coefficient (“MDOAC”)
  Independent of specific chromophore concentration (generalize across systems, applications)
  Strongly dependent on target size, depth, absorption strength, tissue attenuation
MDOAC may be a foundation metric that supports several image quality tests
  Contrast, sensitivity, max visualization depth
  Contrast-detail analysis
  Multispectral biomarker measurement accuracy limitations, e.g., SO2

## Slide 21: Consensus Set of Image Quality Metrics
IPASC should establish a set of image quality metrics supporting test methods
Extensive literature, standards available for mature imaging modalities
Many potentially valid approaches for each performance characteristic
Quantitative resolution test (US)
Qualitative resolution test (CT)

## Slide 22: Consensus List of Terms and Definitions
Spatial measurement accuracy (1D, 2D, or 3D)
MRI (ACR)
Ultrasound (IEC 69391-2)

## Slide 23: Image Quality Metrics
 | Ultrasound | CT | MRI
Sharpness / Spatial Resolution | X | X | X
Section Thickness | X | X | X
Spatial Measurement Accuracy | X | X | X
Distortion | X |  | X
Uniformity | X | X | X
Visualization Depth | X |  | 
Contrast: low contrast resolution, contrast-detail analysis | X | X | X
Artifacts | X |  | X
Material Property Meas. Accuracy | X | X | 
Sensitivity/Noise/SNR | X | X | X
Dynamic Range | X |  | 
Linearity | X |  | 

## Slide 24: Summary and Next Steps
A comprehensive list of IPASC terms and definitions has been developed
  Hold consensus vote to adopt IPASC terms and definitions document
Establish consensus set of image quality metrics to support/guide all theme group activities
  Provide literature review of medical image quality test methods and standards
  Survey IPASC members to prioritize/rank candidate metrics
  Hold discussions to identify, rank utility of image quality metrics
Develop draft multi-site study protocol/SOP

## Slide 25: Afternoon discussion
Which image quality metrics do you think are valuable for PAI system characterization?
Which metrics do you use to evaluate or compare PAI or other imaging devices? How do you measure/compute these metrics?
Are there other metrics that should be considered?
Do you see other challenges that should be addressed?

## Slide 26: Thanks for your attention!